USER MANUAL

LED Curing Lights

[\�'SS3'S	5\AJ!!=T

CE	coxo®

Table of Contents

- Product description	1

- Safety notes	2

- Product List	5

- Components	6

- Presentation	7

- Instruction	10

- Maintenance	13

- Cleaning, Disinfection and Sterilization	14

- Troubleshooting	17

- Technical Specifications	18

- Operating environment.Storage and Transportation conditions•� 9

- Guarantee	19

- Recycling and Disposal	20

- Symbol Identification	20

- Electromagnetic Compatibility Precaution	21

## Product description

The instrument is a cordless pen-style, LED light polymerization instrument for use by dental professionals in dental offices or dental laboratories.

- Small size and lightweight ergonomic design.

- Compact cordless design.

- Individually adjustable Cure tips, rotatable by 360°.

- Cure tip design providing excellent intra-oral access.

##### Contraindications

- The instrument is contraindicated for use in patients prone to photobiological reactions (including patients with solar urticaria or erythropoietic protoporphyria) or those currently undergoing treatment with photosensitizing pharmaceuticals.

- Use with caution for pregnant women, children, and people who have had retina.

- This instrument is not suitable for doctors or patients who have a pacemaker (or other electrical instrument) and are warned not to use small appliances (such as electric shavers, hair dryers, etc.), and hemophiliacs.

##### Intended use

- Used to cure light-curable materials.

- The instrument must only be used in hospital environments, clinics or dental offices, by qualified practitioners.

# Safety notes

Be aware of the following general safety notes and the special safety notes in other chapters of these Instructions for Use.

.ffi.WARNING:

This alerts the user of possibility of extremely serious injury or complete destruction of the instrument as well as other property damage including the possibility of fire.

.ffi. CAUTION:

This alerts the user of possibility of minor or moderate injury or damage to the instrument.

NOTE:

Informs the user of important points concerning operation or the risk of instrument damage.

##### Handpiece

- Don't see the light with your eyes without protection; it's harmful to your eyes. In order to avoid the eyes getting in touch with the light directly, please assemble the eye protector.

- Do not use a damaged instrument; e .g. if among other things, light source exit glass of Cure tip is scratched, broken or missing.

- Only authorized technicians should repair the handpiece.

##### Charging base

- Low voltages are present inside the charging base. Use only under dry conditions. Do not use if the charging base or handpiece is wet. Avoid short circuit between contact pads on the

charging base. Only authorized technicians should repair the charging base.

- Do not use for voltages different from the range indicated on the charging base and power adapter.

- Always make sure handpiece and cure tips is completely reprocessed and thoroughly dry before inserting it into the charging base.

#### Precautions

This product is intended to be used only as specifically outlined in these Instructions for Use.

- Do not use an Instrument which has not been properly reprocessed. Protection sleeves are intended for single use only in order to prevent cross-contamination.

- Never aim the light directly at unprotected soft tissues, as this may cause injury or irritation. Do not aim the light at eyes. Light reflected from the tooth surface may also injure eyes. Use silicone shields supplied with cure tip or eye protector unit.

- Limit the action of the light to the area being treated.

- All dental curing lights cause a certain degree of heat development. Extended operation in areas near the pulp or soft tissues may result in severe damage. So do not cure for more than 20 seconds at a time.

- During heavy use, it is possible for cure tip, which is an applied part, to reach up to 45.5° C. There should be no adverse effects resulting from short-term contact with intact skin or mucous membrane.

- Use only Manufacturer supplied cure tip, adapter and charging base. Use of any accessories other than specified in these Directions for Use may result in damage to the Instrument and its components.

- Sterilizing the Instrument, components and accessories will cause component damage and may cause bodily injury. The silicone shields may be autoclaved.

- Failure to follow recommendations for environmental operating conditions could result in injury to patients or users.

- Inspect instrument before each use for worn, loose or damaged parts.

- There are no user serviceable parts except the 0-ring attached to the coupling end of the light tips. Opening any of the components may result in unsafe operation and will void the warranty.

- This instrument must not be used in the presence of a flammable anesthetic gas mixed with air, oxygen or nitrous oxide.

- User should not touch the patient and accessible charging base contacts simultaneously.

- Wear suitable protective eyewear, mask, clothing and gloves. Protective eyewear is recommended for patients.

- Do not spray disinfectant or other fluid directly onto the cure tips, battery, charging base, adapter.

- Prevent liquids from entering handpiece, charging base, cure tip.

- Do not place the system on or next to a radiator or other heat source. Excessive heat may damage the system's electronics.

- Caries detection mode cannot be used for curing of repair resin.

- Unstable voltage and electromagnetic field will interfere with the normal operation of the instrument.

## Product List

�	c,rntip

Charging base

Adapter	eye protector unit

Handpiece  Single use protection sleeve Silicone shield

.ffi. CAUTION:

If the accessories are damaged, please purchase original accessories and replace and use them according to the instructions.

## Components

##### Handpiece

Cure tip

[\

ON /OFF B,noo

LCD display

Mode Button

Handpiece

##### LCD display

Pulse mode

Caries detection mode Bright light mode

-Cure times

CUR CHK TUR

,.--.,,.-.,

- -

Charging base

Adapter jack

Handpiece charger

Handpiece charging indicator

# Presentation

### Connection

- Cure tip: Connect Cure tip to handpiece by firmly pressing Cure tip into handpiece opening while slightly rotating at the same time

NOTE:

- Cure tip can be rotated 360°.

- In standby state, Power/Start button unavailable.

hen cure tip is not connected or the connection is not in place, and E3 will appear accompanying by two beeps.

- Silicone shield: Hold the Cure tip over the opening of the silicone shield and turn the Cure tip against the silicone shield by 90° into its final position.

Back	Front

NOTE:

- Make sure silicone shield is securely fixed to instrument to avoid accidental aspiration (press silicone shield firmly into correct position).

- Make sure light aperture is not covered by the silicone shield.

- Silicone shield: Attach silicone shields supplied with the handpiece in combination with single use protection sleeve. Hold the Cure tip over the opening of the silicone shield and turn the Cure tip against the silicone shield by 90° into its final position.

NOTE:

- Make sure silicone shield is securely fixed to instrument to avoid accidental aspiration (press silicone shield firmly into correct position).

- Make sure light aperture is not covered by the silicone shield.

- Eye protector unit: Put eye protector unit on cure tip as below.

�

### Battery charge

- Insert handpiece into the charging base.

- Connect the main power supply and charging base via adapter to

charge.

.... DD

•A 4-hour charging is needed generally. The battery charging light will illuminate solid yellow, indicating battery is charging and will show constant green when fully charged.

�Yellow	�	Green

( yellow light) Charging	( green light) Fully charged

NOTE:

- An initial 5-hour burn-in charge is needed prior to first use.

- When the battery is low, the mode icons flash.

- When the battery is very low, the LCD display displays E1,and the instrument automatically shuts down after three beeps.

.ffiWARNING:

- Only use original adapter.

- Do not charge in humid places.

- The battery is not removable.

## Instruction

  - Power on/off

- Short press Power/Start Button to turn it on and it enters stand by state, long press Power/Start Button to turn it off

NOTE

In standby mode, instrument will automatically shut down after three minutes of inactivity.

## Operation

The instrument integrates three modes namely:  ulse mode, caries detection mode and bright light mode.

| Pulse Mode | CUR | Used for curing of repair resin Times: 20 seconds |
| --- | --- | --- |
| Caries Detection Mode | CHK | Used for caries detection Times: 60 seconds |
| Bright Light Mode | TUR | Used for curing bracket binder in orthodontic treatment Times: 3 seconds Not suitable for curing of repair resin |
| Power/Start button |  | Turn the instrumenton/off, start/ stop it |

Mode button	0	Change mode

| Operating Time |  | Demonstrate operating time |
| --- | --- | --- |

Choose mode

      - The mode after power on is the mode when the power was last turned off.

      - Press and hold the mode button for 1 second to switch between pulse, caries detection and bright light modes. The mode icon lights up when switching to caries detection.

##### Start working

.ffi. WARNING:

- Doctors, patients or other personnel must wear goggles during surgery, otherwise it may cause damage to the eyes.

- If you feel the handle becomes hot, stop working until the handle cools down completely.

NOTE

•• During curing, Mode button unavailable.

•• When temperature of microelectronic chip reaches 60°C, E2 will

- appear and instrument will automatically stop

•• It is forbidden for the light source rod to come into contact with

- the patient's mouth or the light emitted by the light source rod to be directed toward the mouth.

- Pulse/Bright light Mode

      - In state by state, short press Power/Start Button to start working. During working, times/cycle(s) counts down until it reaches 0.

      - There will a beep after 5 seconds or an orthodontic treatment.With two beeps after cure times/ orthodontic cycle(s) is

finished.

        - During working, press Power/Start Button to suspension of work.

ill WARNING:

Bright light mode cannot be used for resin curing, otherwise it will cause harm to the patient.

## Caries detection mode

Keep a distance of 5-10mm between the light source rod and the teeth.

In standby mode, short press the power/start button to start the work. During working, the time/cycle displayed in a countdown. When O is displayed, the work ends and is accompanied by two beeps.

Stop using before the end of the cycle, press the power/start button again to stop the work. During the detection process, the fluorescence color is:

  - (Green) Healthy teeth • (Orange) Caries

ill WARNING:

- To ensure accurate detection results, please turn off the external power supply.

- When deep caries are removed, there is fluorescence or a brown color near the pulp.

- After deep caries have been removed, it is recommended to

check your mouth again.

12

## Maintenance

- Before each use, check whether the disposable protective cover, silicone hood, light source rod and other parts are damaged. If so, please stop using it or replace the parts immediately;

- Check whether the handle is damaged before each use. If so, please stop using it immediately and contact our company or authorized dealer for help;

- To prevent cross-infection, it is recommended to use disposable protective covers. Even if a disposable protective cover is used, the light source rod must be cleaned and disinfected before or after use;

- If you want to clean or disinfect, please wipe the surface with water or a disinfectant with an ethanol content of 70%-80%. Do not soak;

- During the cleaning process, prevent liquid or other foreign objects from entering the handle, charging base and light source rod;

- When cleaning the charging base, please disconnect the power supply;

- After each use, please check whether there is any resin residue on the mirror surface of the light source rod, so as not to affect the life of the light source rod or the curing effect.

## Cleaning, Disinfection and Sterilization

| Instrument: | Handpiece, adapter, Cure tip, eye protector unit and charging base, silicone shield. The procedure for cleaning, disinfection and sterilization applies only to silicone shield. |
| --- | --- |
| ADVICE: | Reprocessing procedures have only limited implications to a surgical instrument. The limitation of the numbers of reprocessing procedures is therefore determined by the function / wear of the instrument. There is no limit of maximum allowable reprocessing cycles. The instrument should no longer be reused in case of signs of material degradation. In case of damage the instrument should be reprocessed before sending back to the manufacturer for repair. |

I Reprocessing Instructions

| Preparation at the Point of Use: | Remove silicone shield form handpiece. Remove gross soiling of the instrument with cold water (<40°C) immediately after use. Don't use a fixating detergent or hot water (>40°C) as this can cause the fixation of residuals which may influence the result of the reprocessing process. Store the instruments in a humid surrounding. |
| --- | --- |
| Transportation: | Safe storage and transportation to the reprocessing area to avoid any damage and contamination to the environment. |
| Preparation for Decontamination: | The instruments must be reprocessed in a disassembled state. Only silicone shield can be cleaned and disinfected with automated methods and sterilized with steam sterilization process. Do not sterilize handpiece, adapter, Cure tip, eye protector unit and charging base. The handpiece, adapter, Cure tip, eye protector unit and charging base cannot be cleaned and disinfected in a washer/disinfector. For these parts, only general wipe decontamination is possible! |

| Decontamination of other parts than silicone shield: | After operation.take out Handpiece, adapter, Cure tip eye protector unit and charging base on the work bench. Soak a soft cloth completely with distilled water or deionized water, surfaces of these components, until the surface of the parts the components is visually clean. For decontamination, soak a dry soft cloth with 75% alcohol Wipe all surfaces of Handpiece adapter, Cure tip with the wet soft cloth for about 3 minutes. Please follow the instructions of manufacturer of disinfectant swipe the surface of the component with a dry soft lint-free cloth. |
| --- | --- |
| Pre-Cleaning: | Following instruction is only relevant for silicone shield. Not use automated cleaning, disinfection and sterilization for other parts than silicone shield in this system! Do a manual pre-cleaning, until the instruments are visually clean. Submerge the instruments in a cleaning with cold tap water for at least 10 seconds. Clean the surface with a soft bristol brush. |
| Cleaning: | Regarding cleaning/disinfection, rinsing and drying, it is to distinguish between manual and automated reprocessing methods. Preference is to be given to automated reprocessing methods, especially due to the better standardizing potential and industrial safety. Automated Cleaning: Use a washer-disinfector meeting the requirements of the ISO 15883 series. Put the instrument into the machine on a tray. Connect the instrument with the WD by using suitable adapter and start the program: - 4 min pre-washing with cold water (<40°C) - emptying 5min washing with a mild alkaline cleaner at 55°C - emptying 3 min neutralising with warm water (>40°C) |

•

|  | emptying - 5 min intermediate rinsing with warm water (>40°C) - Emptying The automated cleaning processes have been validated by using . 5% neodisher MediClean forte (Dr Weigert) Ace to en ISO 17664 no manual reprocessing methods are requieror these devices. If a manual reprocessing method has to be used, please validate it prior to use. |
| --- | --- |
| Disinfection: | Automated Disinfection Automated under consideration of national requirements in regards to AO-Value (see EN 15883). A disinfection cycle of 5 min disinfection at 93°C has been validated for the instrument to achieve an AO value of 3000. |
| Drying: | Automated Drying: Drying of outside of instrument through drying cycle of washer/disinfector. If needed, additional manual drying can be performed through lint free towel. lnsuffate cavities of instruments by using sterile compressed a r. |
| Functional Testing, Maintenance: | Visual inspection for cleanliness of the instruments and reassembling. Functional testing according to instructions of use. If necessary, perform reprocessing process again until instrument is visibly clean. Defective accessories should be immediately discarded. The defects include: plastic deformation and corrosion Maintenance is not required. Instruments oil must not be used |
| Packaging: | Pack the instruments in an appropriate packaging material for sterilization. The packaging material and |
| Sterilization: | Sterilization of instruments by applying a fractionated pre-vacuum steam sterilization process (according to EN 285 / EN 13060 / EN ISO 17665) under consideration th respective country requirements. M in at 134° C. In EU, 5 min at 134 |

|  | Maximal sterilization temperature: 137°C Flash sterilization is not allowed on lumen instruments. |
| --- | --- |
| Storage: | Storage of sterilized instruments in a dry, clean and dust free environment at modest temperatures refer to label and instructions for use. |
| Reprocessing The above-mentioned reprocessing process (cleaning, validation study disinfection sterilization) has been successfully information validated. |  |
| Additional Instructions: None |  |
| It is the duty of the user to ensure that the reprocessing processes including resources, materials and personnel are capable to reach the required results. State of the art and often national law requiring these processes and included resources to be validated and maintained properly. |  |

.

## Troubleshooting

| Malfunction Cause |  | Remedy |
| --- | --- | --- |
| Failure to turn power on | Battery is exhausted | Charge battery |
|  | Battery protection | Charge battery |
|  | Battery is damaged | Contact dealer |
| Weaker light | Resin remaining on Light source exit glass surface | Remove the resin |
|  | ow battery | Charge battery |
|  | Surface of light source exit glass is worn | Contact dealer |
|  | LED is damaged | Contact dealer |
| Failure to charge | Poor power connection | Check the connection between adapter and charging base |
|  | Wrong specification of adapter | Check the specifications |
|  | Adapter is damaged | Contact dealer |

| Shorter operation | battery aging | Contact dealer |
| --- | --- | --- |

The parts of the instrument should be repaired by designated professionals If problem cannot be solved, please contact local dealer or our company.

## Technical Specifications

| Adapter | lnput:AC100-240V/50 /60Hz |
| --- | --- |
|  | Output DC 5V, 1.5A |
| Input power | 10VA |
| Battery performance | 3.7V1400mA/h |
| Luminous power | More than 1000 mw/cm2 |
| Output peak wavelength range | 385 nm - 515 nm |
| Output peak center wavelength | 460nm |
| 385nm-515nm (Blue-ray) Optical radiation wavelength range | >200 mw/cm2 |
| 2 | <200mW/cm2 |
|  | <100mW/cm2 |
| Effective area of the cure tip | 78.5mm2 |
| Operation mode | Non-continuous |
| Degree of Protection(IEC 60529) | IPX0 |
| Protection against Electric Shock | Type B applied part |
| Classified of protection against Electric Shock | Class II ( Adapter) |
| Handpiece dimensions | L: 22.5cm;W : 2.67cm |
| Handpiece weight | 84.7 grams |

J  Applied part	J  Cure tip, silicone shield

# Operating environment Storage and Transportation conditions

# Guarantee

- H

|  | Direct current |  | Sterilizable in a steam sterilizer at 134°C |
| --- | --- | --- | --- |
| - | Products comply with WEEE directive |  |  |

## Electromagnetic Compatibility Precaution

This information is required by the 4th edition of IEC 60601-1-2.

  - The Unit needs special precautions regarding EMC and needs to be installed and put into service according to the EMC information provided in this manual.

  - Portable and mobile RF communications equipment can affect the Unit.

  - The use of accessories, transducers and cables other than those specified by manufacturer,may result in increased emissions or decreased immunity of the Unit.

  - The Unit should not be used adjacent to or stacked with other equipment and that if adjacent or stacked use is necessary, the Unit should be observed to verify normal operation in the configuration in which it will be used.

  - Per IEC 60601-1-2, no additional environmental operating conditions are required for normal use.

| Guidance and manufacture's declaration - electromagnetic emission |  |  |  |  |
| --- | --- | --- | --- | --- |
| The instrument is intended for use in the electromagnetic environment specified below. The customer or the user of the instrument should assure that it is used in such an environment. |  |  |  |  |
| Emission test | I | ComplianceI | Electromagnetic guidance | environment - |
| RF emissions CISPR 11 |  | Group 1 | The instrument use RF energy only for its internal function. Therefore, its RF emissions are very low and are not likely to cause any interference in nearby electronic equipment. |  |
| RFemission CISPR 11 |  | Class B | The instrument is suitable for use in all establishments, including domestic establishments directly connected to the public low - voltage power supply network with specific requirement. |  |
| Harmonic emissions IEC 61000-3-2 |  | Class A |  |  |
| Voltage fluctuations/flicker emissionslEC 61000-3-3 |  | Complies |  |  |

| Guidance and manufacture's declaration - electromagnetic immunity |  |  |  |
| --- | --- | --- | --- |
| The instrument is intended for use in the electromagnetic environment specified below. The customer or the user of instrument should assure that it is used in such an environment. |  |  |  |
| Immunity test I IEC 60601 Compliance level ·_1 Electromagnetic test level I environment - guidance |  |  |  |
| Electrostatic discharge (ESD)IEC 61000-4-2 | ±8 kV contact ±2 kV, ±4kV, ±8kV,±15 kV air | ±8 kV contact ±2 kV, ±4kV, ±8kV,±15 kV air | Floors should be wood, concrete or ceramic tile. If floor are covered with synthetic material, the relative humidity should be at least 30%. |
| Electrical fast transient/burst IEC61000-4-4 | ±2kV for power supply lines ±1 kV for Input/output lines | ±2kV for power supply lines ±1 kV for Input/output lines | Mains power quality should be that of atypical commercial or hospital environment. |
| Surge IEC 61000-4-5 | ±0.5kV, ±1 kV line to line ± 0.5 kV, ±1 kV,± 2 kV line to ground | ±0.SkV, ±1 kV line to line | Mains power quality should be that of a typical commercial or hospital environment. |

|  | communication instrument(Refer to table 9of IEC 60601-1-2:2014) | 385MHz -5785MHz Test specifications forENCLOSU REPORT IMMUNITY to RF wireless communication instrument( Refer to table 9 of IEC60601-1-2:2014) | Where P is the maximum output power rating of the transmitter in watts (W) according to the transmitter manufacturer, and "d" is the recommended separation distance in meters (m). Field strengths from fixed RF transmitters, are determined by an electromagnetic site survey should be less than the compliance level in each frequency range 'Interference may occur in the vicinity of instrument marked with the following symbol: (('i')) |
| --- | --- | --- | --- |
| NOTE 1 At 80 MHz and 800 MHz, the higher frequency range applies. NOTE 2 These guidelines may not apply in all situations. Electromagnetic Propagation is affected by absorption and reflection from structures, objects and people. |  |  |  |
| a Field strengths from fixed transmitters, such as base stations for radio (cellular/ cordless) telephones and land mobile radios, amateur radio, AM and FM radio broadcast and TV broadcast cannot be predicted theoretically with accuracy. To assess the electromagnetic environment due to fixed RF transmitters, an electromagnetic site survey should be considered. If the measured field strength in the location in which the instrument is used exceeds the applicable RF compliance level above, the instrument should be observed to verify normal operation. If abnormal performance is observed, additional measures may be necessary, such as re-orienting or relocating the instrument 'Over the frequency range 150 kHz to 80 MHz, field strengths should be less than3 V/m. |  |  |  |

The instrument is intended for use in an electromagnetic environment in which radiated RF disturbances are controlled. The customer or the user of the instrument can help prevent electromagnetic interference by maintaining a minimum distance between portable and mobile RF communications equipment (transmitters) and the instrument as recommended below, according to the maximum output power of the communications equipment.

Recommended separation distances between portable and mobile RF communications equipment and the instrument.

| Rated maximum output power of transmitter(W) | Separation distance according to frequency of transmitter |  |  |
| --- | --- | --- | --- |
|  | 150 kHz to 80 MHz d=1.2xP1l2 | 80 Mhz to 800 MHZ d=1.2xP1l2 | 80 MHz to 800 Mhz d=2.3xP1l2 |
| 0.01 | 0.12 | 0.12 | 0.23 |
| 0.1 | 0.38 | 0.38 | 0.73 |
|  | 1.2 | 1.2 | 2.3 |
| 10 | 3.8 | 3.8 | 7.3 |
| 100 | 12 | 12 | 23 |
| For transmitters rated at a maximum output power not listed above, the recommended separation distanced in meters (m) can be estimated using the equation applicable to the frequency of the transmitter, where P is the maximum output power rating of the transmitter in watts (W) according to the transmitter manufacturer. NOTE1: At 80 MHz and 800 MHz, the separation distance for the higher frequency range applies. NOTE2: These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects and people. |  |  |  |

Foshan COXO Medical Instrument Co., Ltd. No.17,Guangming Ave.,New Light Source Industrial Base, Nanhai National High-ech Zone, Foshan 528226, Guangdong P.R. China

!ECI REP! Lotus NL B.V.

Koningin Julianaplein 10, 1e Verd, 2595AA, The Hague, Netherlands.

E-mail:peter@lotusnl.com

Ver: 1.0 Date: 20240513